SOP

3D Printers

Form 3

Training is required before using this equipment

Reservations	are required	when using this equipment

# Reference Sheet

Approved Materials:

Only Formlabs brand resin may be used on this printer

Build Volume: Layer Resolution (min.):

5.7” x 5.7” x 7.3” (14.5 x 14.5 x 18.5 cm)

Layer Thickness:

25-300 microns (dependent on resin type)

Machine Accessories:

| • | Form Wash | • | Form Cure |
| --- | --- | --- | --- |

Important Information:

Resin can be a skin irritant, care should be taken to not make contact with resin

Workflow:

Download or Design

Process in Preform

Setup Machine

3D Print Object

Remove finished part(s) from printer

Clean part using Form Wash

Cure part using Form Cure

Weigh and pay

Remove supports

Software:

Formlabs Preform

Tinkercad

Fusion 360

Solidworks

# Machine layout - Front

- Resin Cartridge

- Status Light

- Cover

- Build Platform

- Resin Tank

- Touchscreen

- Status Light

# Removing Material Tank and Cartridge

- To change the material, begin by closing the vent on the top of the cartridge located in the back of the printer.

- Lift the cartridge up and out of the machine.  Store the cartridge with the other resin cartridges.

- Pull the tank away from the back of the machine.

- Lift up with both hands to free the tank from the machine.  Cover the extra tank to protect it from dust and debris while it is in storage.  Store it with the other materials.

Wiper

Arm

Resin Cartridge Location

Image 1 – Inserting Material Tank

# Installing Material Tank

- Select the new material and remove the temporary cover.  Insert the tank and press the tank toward the back of the printer to lock it into place.  The tank will click into place (image 1).

- Slide the resin cartridge into place at the back of the machine.  Open the vent on top of the resin cartridge (image 2).

- Insert a build plate by lifting the central locking arm upward, sliding the build plate onto the arm, and closing the lock (image 3).

Image 2 – Inserting Resin Cartridge

Image 3 – Inserting Build Platform

# Process in FormLabs Preform

- Open Preform.  A blank print area will appear.  Note that the build volume is upside down in this program.

- To add the file to the build platform select File -> Open.  You can also drag and drop your file into the build volume.

- To move the view around, click and hold the middle mouse button while moving the mouse to pan.  Click and hold the right mouse button while moving the mouse to rotate. Scrolling up will zoom in on the cursor and scrolling down will zoom out of the cursor.

- Clicking and dragging on the object lets you move it around the build plate. Pressing delete when the model is selected will

remove it from the build plate.  Clicking and dragging on the spherical icon on the object allows you to rotate the object.

- There are 6 icons in the upper left hand corner of the Preform interface.

  - The first icon is the magic wand which will automatically rotate the object to optimize it for SLA printing and add supports where needed.

  - The second icon allows the user to manually scale the object.

  - The third icon allows the user to manually rotate the object.

  - The fourth button allows the user to alter the support structure of the object.

  - The fifth icon allows the user to manually place all objects on the print bed.  This button can also automatically pack all models to be a certain spacing apart (specified by the user).

  - The sixth icon is the Print button, which will send the user’s print to the printer.

5

. The

right

of the screen shows information on the material and

layer thickness, the estimated volume and time, as well as the

“Printability” of the ob

ject.

6

. Adjust the material or

layer thickness by clicking on the

printer’s name under “Printer.”

Note that not all materials are

able to print in all thicknesses.

7

. Check the “Printability” of the object before sending it to

the

printer. Click on the

Red X or the Green Check Mark to view the

errors Preform anticipates.  To correct errors, try adjusting the

support density or size until Preform shows the part is printable.

. To send the part to the printer select the sixth icon in the upper

8

left

-

hand

corner “Print” and select

“Upload Job” in the bottom

right of the pop

-

up

.  The file will then be sent to the printer

interface.

.

10

To start the pr

int,

tap “Print

” on the printer’s touch screen.

11

.

At this point, the resin will begin to heat up and

will start

printing.

Print Button on Printer Touch

Screen

# Cleaning Part

- When the Form 3 is finished with the print job, it will retract the build tray to the top of the machine.  Remove the entire build tray by raising the lock lever and sliding the whole tray out.

- Open Form Wash by moving the dial and selecting “Open.”

- If needed adjust the time for the rinse cycle.  10 minutes is sufficient for most parts.

- Insert the whole build platform to the top of Form Wash and on screen select “Start.”

- When the rinse cycle is complete Form Wash will automatically raise.  Allow the part to drip dry for several minutes before retrieving it.

- Once the part is completely dry, attach the build platform to the holding jig and use the included tools to remove the print from the build platform.  Note that parts are still somewhat fragile in this state. Parts must be cured to achieve full strength.

# Curing Part

- Prints need additional curing after printing in order to achieve their full strength.  Consult the included chart to see curing time and temperature. Adjust time and temperature on Form Cure as needed.

- Open the door on Form Cure and place the printed item in the middle of the table.  Close the door and select “Start” on the control screen.

- Form Cure will preheat to the selected temperature and then begin the curing process.  During this time, the part will be exposed to ultraviolet light and it will fully harden.

## Cure Time and Temperature Settings

| Material | Time (min) Temp (°C) |
| --- | --- |
| Clear V3 | 15 60 |

White V4  	30 	60

Grey V4  	30 	60

Black V4  	30 	60

Color V1  	30 	60

Tough V5  	60 	60

Durable V2  	60 	60

Flexible V3  	15 	60

High Temp V1  	30 	60

Rigid V1 	15 	80

Grey Pro V1  	15 	80

Dental SG V1  	30 	60

Dental LT Clear V1  	20 	80

Model V2  	30 	60

Castable V2  	240 	60

Ceramic V1  	N/A  	N/A

Castable Wax  	N/A  	N/A

# Removing Supports

Parts should be weighed before supports are removed

- After curing is complete supports may be removed.  Wear safety glasses as small parts will move unexpectedly.

- Using diagonal clippers remove small pieces of supports at a time until all support pieces are gone.

- For highly cosmetic parts a hot knife may be used.

# Troubleshooting

If parts come loose and fall off the build platform during printing they must be removed.  Alert NIS Staff to filter the resin.